## Příloha č. 1 Smlouvy
Technical conditions and Technical specification of the offered system |  | Appendix No. 1 of the Contract
Programmable puller of glass microelectrodes
TECHNICAL CONDITIONS technical specification required by the customer |  | TECHNICAL SPECIFICATION OF THE OFFERED SYSTEM technical specification offered by the supplier
Description of the parameter (The offer must meet all the requirements and specification of the parameters listed below. For parameters defined by a minimum or maximum level or a range of values, the offer must meet at least the specified required level.) | Value required by the customer | (If the customer only requires the supplier to indicate whether the relevant required parameter is met, then the supplier indicates YES if the parameter is met. If the customer requires a more detailed description or determination of the specification for a technical parameter, then the supplier is obliged to provide a more detailed description, a list of properties, a specific data, or a range of values of the offered goods.)
(It must be clear from the data provided by the supplier that the goods offered by the tenderer meet the minimum technical requirements set by the customer - the supplier shall indicate the fulfillment of the required parameter in a verifiable manner, e.g. by stating specific values, or by specific reference to technical sheets, drawings, etc.)
General requirements of the customer on the device: Delivery of a new device that has not been used before and was not refurbished is required. The device has to allow the pulling of glass microelectrodes usable for both patch clamp and sharp electrode techniques. A high level of preciseness and reproducibility of the pulling is required to obtain the same diameter of the orifice of the microelectrodes pulled with the same program and parameters. Automatic attachment of microcapillary and automatic heat-polishing of the tip of microelectrode as a part of the program have to be possible.
Programmable puller of glass microelectrodes |  | Model - type/production designation | Producer
|  | will be filled in by the supplier | will be filled in by the supplier
Number of pieces: 1 |  | Value offered by the supplier
General requirement | Delivery of a new device that have not been used before and was not refurbished is required | supplier must confirm required fact - quote YES or NO to indicate by fact
Programs - attachment of the microcapillary | automatic attachment of the pulled microcapillary | supplier must confirm required fact - quote YES or NO to indicate by fact
Programs - number | it should be possible to define and save at least 60 different programs for the microelectrode pulling | supplier must confirm required fact - quote YES or NO to indicate by fact
Programs - stages | one- and two-stage pulling programs have to be available, the stages have to be defined independently, at least 6 preliminary pulls should be available | supplier must confirm required fact - quote by the fact
Programs - parameters | both heating and pulling force and time have to be adjustable to adapt the program for pulling of the required type of microelectrodes | supplier must confirm required fact - quote YES or NO to indicate by fact
Programs - types of microelectrodes | has to enable the pulling of both patch microelectrodes (a tip between at least 1 and 60 µm) and sharp microelectrodes (maximal resistance - at least 90 MΩ) | supplier must confirm required fact - quote by the fact
Programs - microelectrodes | has to allow automatic compensation for variation in the thickness of the glass capillary | supplier must confirm required fact - quote YES or NO to indicate by fact
Programs - microelectrodes | has to enable automatic heat-polishing of the microelectrode tip as a part of the program | supplier must confirm required fact - quote YES or NO to indicate by fact
Compatibility with microelectrodes | has to be compatible with the borosilicate glass microcapillaries of the outer diameter of 1.5 mm and the length of 7.5 -8.0 cm | supplier must confirm required fact - quote by the fact
No need of periodical calibration | Yes | supplier must confirm required fact - quote YES or NO to indicate by fact
User guide (in Czech or English) | Yes | supplier must confirm required fact - quote YES or NO to indicate by fact
Power supply | 230 V, 50-60 Hz | supplier must confirm required fact - quote YES or NO to indicate by fact
Maximal dimensions (width 60 cm x depth 40 cm x height 40 cm) | Yes + report the values | supplier must confirm required fact - quote by the fact
Weigth between 25 and 35 kg | Yes + report the value | supplier must confirm required fact - quote by the fact
Warranty period | at least 24 months | supplier must confirm required fact - quote by the fact