Technical Specifications

PURPOSE AND SUBJECT OF THE PUBLIC CONTRACT

The purpose of the delivery is to obtain a system for operant IV self-administration.

The subject of delivery is new and fully functional equipment including complete installation, commissioning and training of personnel at the delivery site.

The bidder's bid must meet all the requirements and parameters specified by the Contracting Authority below. In the case of parameters defined by a minimum or maximum level or range of values, the tenderer's tender must meet at least the specified required level.

MINIMUM TECHNICAL PARAMETERS

| Parameter | Parameter offered by the supplier |
| --- | --- |
| Producer | (Supplier to specify manufacturer) |
| Type/Model | (Supplier to specify type and model) |

| Minimum required technical parameters | Technical parameters offered by the supplier* |
| --- | --- |
| 16 operant chambers | Supplier shall specify YES/ NO |
| Laptop to control the system at least 15 inch diagonal | Supplier shall specify YES/ NO |
| Data acquisition software for 1 Laptop | Supplier shall specify YES/ NO |
| SW for data analysis, unlimited license | Supplier shall specify YES/ NO Supplier shall specify YES/ NO |
| Starter pack of system compatible consumables | Supplier shall specify YES/ NO |

| Specification of operant chambers (pcs per chamber) |  |
| --- | --- |
| 1 sound-attenuating cubicle | Supplier shall specify YES/ NO |
| 1 fan (minimum) | Supplier shall specify YES/ NO |
| 3 retractable levers | Supplier shall specify YES/ NO |
| 3 nose-poke operands | Supplier shall specify YES/ NO |
| 3 stimulus “cue” lights (white) | Supplier shall specify YES/ NO |
| Tone generator, minimum 1 fixed tone generator | Supplier shall specify YES/ NO |
| IR beams for tracking of locomotor activity during session | Supplier shall specify YES/ NO |
| IR light placed around and under the lid of the chamber for experiments performed in the dark | Supplier shall specify YES/ NO |
| A bracket to place the camera for monitoring the experiment | Supplier shall specify YES/ NO |
| 2 LED house lights | Supplier shall specify YES/ NO |
| LED light source located under the chamber lid for uniform illumination of the chamber | Supplier shall specify YES/ NO |
| The chamber for housing the laboratory rat must be on a sliding shelf for easy access and cleaning of the chamber | Supplier shall specify YES/ NO |
| The sliding shelf must be secured against falling out and at the same time must be easy to slide out. | Supplier shall specify YES/ NO |
| The sliding shelf must have a load capacity of at least 25 kg and must provide sufficient support for working with the chamber. | Supplier shall specify YES/ NO |
| 1 pellet dispenser | Supplier shall specify YES/ NO |
| 1 reward tray with head entry detector | Supplier shall specify YES/ NO |
| 2 floors | Supplier shall specify YES/ NO |
| 2 waste pans | Supplier shall specify YES/ NO |
| 1 infusion pump with fixed speed (allowing 0.1 mL/5 seconds from 10 ml syringe) | Supplier shall specify YES/ NO |
| The infusion pump must be placed on a sliding shelf above the chamber for easy access and handling | Supplier shall specify YES/ NO |
| 2 one-channel stainless steel swivels | Supplier shall specify YES/ NO |
| 1 drug delivery arm | Supplier shall specify YES/ NO |
| The swivel must be placed on a sliding shelf above the chamber for easier access and handling | Supplier shall specify YES/ NO |
| Individual chambers must be equipped with their own control processor | Supplier shall specify YES/ NO |
| Chambers must be able to carry out experiments even in the event of technical malfunctions (non-functioning PC, ...) | Supplier shall specify YES/ NO |
| The software and hardware of individual chambers must work in a complete set of all supplied chambers without mutual interference and data drops. | Supplier shall specify YES/ NO |
| The entire system must be easy to clean – the lid of the chamber must be removable and at the same time it must be possible to lift the walls of the chamber. | Supplier shall specify YES/ NO |

| Starter pack specifications |  |
| --- | --- |
| The starter pack contains consumables fully compatible with the offered system without the need to purchase additional components | Supplier shall specify YES/ NO |
| The starter package includes at least the following items in the following quantities: - Jugular vein catheter for rats – 32 pcs - Intravascular button for rats - 32 pcs - Magnetic tether – 32 pcs - Protective cover for buttons – 32 pcs - Tubing - with a total length of 50 m - Connectors – 32 pcs - Heparin – 32 pcs | Supplier shall specify YES/ NO |

| General requirements |  |
| --- | --- |
| Trolley for stacking cubicles in two layers, with wheels | Supplier shall specify YES/ NO |
| The boxes must be firmly connected to the trolley to ensure maximum safety for the personnel | Supplier shall specify YES/ NO |
| System for safe and easy installation of the infusion pumps (internal arrangement so that the system is enclosed in a box) | Supplier shall specify YES/ NO |
| Trolley with boxes must include all necessary hardware and prewired cables for connecting the system with a limited number of required plugs (pre-wired) | Supplier shall specify YES/ NO |
| Plan for sterilization of the entire system when installed in the SPF user device (H2O2 vapor sterilization is necessary) | Supplier shall specify YES/ NO |
| Individual parts of the system (walls, floors, trays) must be removable for more effective sterilization | Supplier shall specify YES/ NO |

* The supplier shall indicate YES/NO and complete the required information.

The Supplier is obliged to attach his own technical specification or his own description of the equipment to the technical specification given in this table.