## Příloha č. 1 Smlouvy
Technical specification |  |  | Annex no. 1
Multielectrode system - Measuring system for recording of data using the multielectrode array (MEA) technique and their analysis
TECHNICAL CONDITIONS required |  |  | TECHNICAL SPECIFICATION offered
General requirements on the equipment "Multielectrode system ": The measuring system is planned for analysis of electrical properties of excitable tissues, namely of the cardiac and neural tissues/cells. Recording of electrical field potentials will enable to analyse excitability of the tissues/cells, spreading of the electrical activity through the cellular networks etc., both in control conditions and under effects of selected drugs. The acquired data will contribute to better understanding of electrophysiological properties of the studied tissues/cells, both in control and under drug effects, and both in healthy tissue and in pathological samples. We require delivery of a new device that have not been used before and was not refurbished.
The Multielectrode system will be composed of the following interconnected and fully compatible parts: 1) MEA acquisition system with integrated voltage/current stimulator, heating element, filter amplifier, analog-digital converter, and software for data acquisition and evaluation; 2) two-channel temperature controller; 3) peristaltic perfusion system.
Item No. 1 MEA acquisition system |  |  | Model - type/production designation | Producer
|  |  | will be filled by supplier | will be filled by supplier
Number of items: 1 pc |  |  | Value offered by the supplier
MEA system |  | filter amplifier with integrated data acquisition and analog-digital converter, compatible with the headstages, at least two headstages can be attached and used at the same time during the same recording | supplier must confirm required fact - quote YES or NO to indicate by fact
Headstage, at least 2 pieces |  | can be operated in a humid incubator, compatible with MEA chips with at least 60 electrodes (59 measuring ones and a reference one), and enabling visual control of the electrode array using both inverted and upright microscopes | supplier must confirm required fact - quote by the fact
Bandwidth filter |  | yes, adjustable at least between 0.1 Hz and 10 kHz (via software) | supplier must confirm required fact - quote by the fact
Data resolution |  | at least 24 bit | supplier must confirm required fact - quote by the fact
Stimulus generators |  | yes, internal, min. 3 independent, direct access to each electrode for stimulation, control of parameters of the stimulation via the included software | supplier must confirm required fact - quote by the fact
Current stimulation |  | min. in the range +/- 1.0 mA | supplier must confirm required fact - quote by the fact
Voltage stimulation |  | min. in the range +/- 10 V | supplier must confirm required fact - quote by the fact
Sampling rate |  | min. 40 kHz per channel | supplier must confirm required fact - quote by the fact
Real time detection of the signal |  | yes | supplier must confirm required fact - quote YES or NO to indicate by fact
Heating element |  | yes, integrated, regulated via the 2-channel temperature controller (Item No. 2) | supplier must confirm required fact - quote YES or NO to indicate by fact
System for stimulus artifact suppression |  | yes, integrated | supplier must confirm required fact - quote YES or NO to indicate by fact
Fixation for the perfusion system |  | yes, near to the recording MEA interface | supplier must confirm required fact - quote YES or NO to indicate by fact
Connection with computer |  | High Speed USB 3.0 or higher | supplier must confirm required fact - quote YES or NO to indicate by fact
Power supply |  | yes, input voltage: 220-240 V AC | supplier must confirm required fact - quote YES or NO to indicate by fact
Inputs and outputs |  | min. 4 digital outputs, min. 4 digital inputs for synchronization, min. 2 analog inputs, min. 1 stereo output, 1 ground port | supplier must confirm required fact - quote by the fact
Software |  | acquisition and analysis software for elelctrophysiological data from both neural and cardiac tissue/cells (unlimited licence; enabling software filtration, upload of the recorded data, their viewing and subsequent processing, export to common data formats, and both real time and ex post analysis), including maintenance package (updates and support) for at least 2 years | supplier must confirm required fact - quote YES or NO to indicate by fact
Manual |  | yes, manual in Czech or English | supplier must confirm required fact - quote YES or NO to indicate by fact
Safety - CE marking |  | Yes | supplier must confirm required fact - quote YES or NO to indicate by fact
Accessories 1 - Holder for the perfusion cannula |  | yes, at least 2 pieces, each with fixation screw to fix the perfusion cannula; attachable near to the MEA interface to enable perfusion of the cells or tissue sample placed on the MEA | supplier must confirm required fact - quote by the fact
Accessories 2 - Climate chamber |  | yes, at least 2 pieces, keeping stable atmosphere (humid or non humid) around the biological sample | supplier must confirm required fact - quote YES or NO to indicate by fact
Accessories 3 - Signal generator |  | Yes | supplier must confirm required fact - quote YES or NO to indicate by fact
Item No. 2 Two-channel temperature controller |  |  | Model - type/production designation | Producer
|  |  | will be filled by supplier | will be filled by supplier
Number of items: 1 pc |  |  | Value offered by the supplier
Temperature - control range |  | ambient to 45 °C | supplier must confirm required fact - quote by the fact
Accuracy |  | 0.1 °C and better | supplier must confirm required fact - quote by the fact
Resolution |  | 0.1 °C and better | supplier must confirm required fact - quote by the fact
Dimensions (width x depth x height) |  | max. 20 cm x 25 cm x 15 cm | supplier must confirm required fact - quote by the fact
Supply voltage |  | 220 - 240 V AC | supplier must confirm required fact - quote YES or NO to indicate by fact
Quick USB interface |  | yes, enabling adjustment of all settings in real time | supplier must confirm required fact - quote YES or NO to indicate by fact
Software |  | yes, enabling adjustment of all settings in real time | supplier must confirm required fact - quote YES or NO to indicate by fact
Manual |  | yes, manual in Czech or English | supplier must confirm required fact - quote YES or NO to indicate by fact
Item No. 3 Peristaltic perfusion |  |  | Model - type/production designation | Producer
|  |  | will be filled by supplier | will be filled by supplier
Number of items: 1 pc |  |  | Value offered by the supplier
Peristaltic perfusion |  | low pulsation pumps, one inlet and one outlet pump; both pumps have to be adjustable independently | supplier must confirm required fact - quote by the fact
Dimensions (width x depth x height) |  | max. 20 cm x 20 cm x 25 cm | supplier must confirm required fact - quote by the fact
Weight |  | max. 5 kg | supplier must confirm required fact - quote by the fact
Adjustment of the flow |  | via touch screen; rate at least up to 25 ml/min for both the inlet and outlet pump | supplier must confirm required fact - quote by the fact
Bubble trap |  | yes, at least 2, in both the inlet and outlet pumps | supplier must confirm required fact - quote YES or NO to indicate by fact
Motor drives |  | with constant flow rates even at lowest setting, with low electromagnetic emission | supplier must confirm required fact - quote YES or NO to indicate by fact
Set of tubing |  | Yes | supplier must confirm required fact - quote YES or NO to indicate by fact
Supply voltage |  | 220 - 240 V AC | supplier must confirm required fact - quote YES or NO to indicate by fact
The cells marked in this way will be filled in by the supplier.