## List1
"MR copatible LCD screen for CEITEC MU (3rd repetition)" Appendix no 1 - Technical specifications
Type designation of the device
To be completed by supplier
Basic requirements
Contracting authority require purchase 1 piece of the MR compatible LCD screen for visual stimulation during acquisition in 3T MRI scanner.
Compulsory parameters
MR compatible LCD screen will be used as an equipment for visual stimulation of subjects during MRI examination. The equipment must be compatible with 3T scanners, especially Siemens Magnetom Prisma. It means the function of LCD screen will not be affected by MR scanner and the function of MR scnner will not be affected by the LCD screen.
Technical and functional characteristics (The bids must meet all parameters set out below.) | Requested value | Supplier's offer Supplier indicates meeting all the technical and functional requirements; shall be evidenced unambiguously and plausibly, e.g. by providing technical sheets and drawings of the manufacturer.
Manipulation requirements | Stand/trolley for easy manipulation (changing/adjusting position) by single person. | To be completed by supplier
Power supply | 230 V AC 50 Hz | To be completed by supplier
Signal source input | DVI or HDMI | To be completed by supplier
Diagonal | min. 24 ", max 40 " | To be completed by supplier
Minimal physical screen resolution - horizontal | min. 1600 pixels | To be completed by supplier
Minimal physical Screen resolution - vertical | min. 1200 pixels | To be completed by supplier
Requested operation in case of higher resolution than 2K | FullHD at 60Hz, or 1920x1200 at 60Hz | To be completed by supplier
Screen refresh rate for full physical resolution | min. 60 Hz | To be completed by supplier
Contrast ratio | min 1000:1 | To be completed by supplier
Grey-to-grey response time | max. 10 ms | To be completed by supplier
Maximal total weight (including power supply and stand/trolley | max. 55 kg | To be completed by supplier
Origin of the goods (for the import duty) | Country | To be completed by supplier
The seller (tenderer) declares that the delivery will comply with all the aforementioned requirements of the purchaser (contracting authority). If it becomes obvious in the course of the delivery preparation and implementation that it is necessary to deliver additional devices, components or accessories or to perform any additional services or works in order to meet the purchaser's requirements hereunder, the seller hereby undertakes to deliver such devices and to perform such works or to provide such services as a part of the original delivery under the agreement, without any increase of the purchase price (i.e. such deliveries, works or services shall not constitute extra deliveries or extra works).