Technical Specifications

Device description:

High Power Impulse Magnetron Sputtering (HiPIMS) is a pulse plasma technology for coating applications. The applied power is focused into a short pulse (from tens of microseconds to hundreds of microseconds), with duty cycles less than 5%. HiPIMS offers dense plasma formation with a high fraction of ionized species and high ion flux towards the substrate, which improves properties of the deposited thin films.

Aim of the tender is the acquisition of a new HiPIMS power source for magnetron sputtering.

Technical specification:

- Two independent power units – each for one channel

- Minimum average power 5 kW per first unit and 10kW per second unit

- Minimum output current and voltage: 1000 V and 500 A for 5 kW unit; 1000 V and 1000 A for 10 kW unit

- Minimum pulse length 20 us and minimum repetition frequency of 50 Hz

- Synchronization unit allowing synchronization of both HiPIMS power units

- Control software

- 1-year warranty

- Order in English

- Payment in CZK

- Online technical support

- Repairs within 10 weeks

- Delivery within 12 weeks from ordering

- Place of delivery: Masaryk University, Faculty of Science, Department of Physical Electronics, Kotlarska 2, 611 37 Brno, Czech Republic