Technical conditions:

Description of device usage:

System for measurements of interaction of biomolecules in real-time mode. It is able to measure hydrodynamic friction to determine the absolute size and shape of interacting molecules using high frequency dynamic electrical switching mode. It is able to measure real-time kinetics of interactions independently of the size of molecules through changes in the local environment by enhancing 
the fluorescence of molecules. The device includes a temperature-controlled automatic sample dispenser, a temperature controlled biochip environment for the measurement itself, and two independent probing fluorescence channels. The device is fully programmable and computer-controlled; it also includes an automated system for conjugation of biomolecules to DNAs controlled by computer station.

Technical specifications:

| Sensing of interactions |  | Technical parameters offered by the supplier * |
| --- | --- | --- |
| Limit of detection LOD | 10 fM | Supplier will indicate YES / NO |
| Dissociation constant KD | At least in the range from 50 fM to 1 mM KD measurements from kinetics and titration | Supplier will indicate YES / NO |
| Association rate constant kON | At least in the range of 1E3 – 1E8 M-1s-1 | Supplier will indicate YES / NO |
| Dissociation rate constant kOFF | At least in the range of 1E-6 – 1E0 s-1 | Supplier will indicate YES / NO |
| Analysis |  | Technical parameters offered by the supplier * |
| Determination of Molecule Size | Hydrodynamic Protein Diameter (DH, Stokes radii on chip) 1 nm < DH < 14 nm (8 kDa – ~500 kDa) | Supplier will indicate YES / NO |
| Determination of Molecule Shape | Conformational changes (qualitative and quantitative if ΔDH ≥ 0.1 nm) | Supplier will indicate YES / NO |
| Temperature control of measurements | Measurement at constant temperatures Measurement in temperature gradients | Supplier will indicate YES / NO |
| Detection of Melting point – Range | At least in the range of 8°C to 75°C | Supplier will indicate YES / NO |
| Melting point - Ramp speed | ΔT/Δt at least up to 10°C/min | Supplier will indicate YES / NO |
| Melting temperature – Accuracy | At least ΔTM ~ 0.5°C | Supplier will indicate YES / NO |
| Melting curve analysis (TM) & small-molecule induced changes in thermal stability | Measurable | Supplier will indicate YES / NO |
| Multiple binding | Detection of avidity, cooperativity Measurements of two different analytes on the same detection spot via dual-colour detection | Supplier will indicate YES / NO |
| Determination of Thermodynamic Parameters | ΔG, ΔH, ΔS | Supplier will indicate YES / NO |
| Equilibrium binding energies | Van’t Hoff analysis 8°-75°C | Supplier will indicate YES / NO |
| Transition state energies | Eyring analysis 8°-75°C | Supplier will indicate YES / NO |
| Software package for analysis and design of automatic measurements | 2 licences, free upgrade for 2 years | Supplier will indicate YES / NO |
| Sample and Liquid handling |  | Technical parameters offered by the supplier * |
| Number of flow channels | At least 4 flow channels | Supplier will indicate YES / NO |
| Number of electrodes | At least 20 | Supplier will indicate YES / NO |
| Electrode material | Gold | Supplier will indicate YES / NO |
| Electrode diameter | Maximum 120 μm | Supplier will indicate YES / NO |
| Conjugation method | Coupling to DNA oligos | Supplier will indicate YES / NO |
| Flow channel height | Maximum 60 µm | Supplier will indicate YES / NO |
| Minimal sample volume | At least from 10 μL to maximum 50 μl | Supplier will indicate YES / NO |
| Pumps | 2 pumps installed: syringe and peristaltic | Supplier will indicate YES / NO |
| Pump Rate | At least in the range from 1 to 4000 μL / min | Supplier will indicate YES / NO |
| Pump Capacity | At least 5 mL (syringe) and NA (peristaltic) | Supplier will indicate YES / NO |
| Temperature Range (flow channel) | 8°C – 70°C (low flow rates) 10°C – 40°C | Supplier will indicate YES / NO |
| Ramp speed (flow channel) | up to 10°C / min | Supplier will indicate YES / NO |
| Autosampler capacity | 36 × 1.5 mL vials & 12 × 10.0 mL vials or 96 well plate & 6 x 10.0 mL & 6 x 1.5 mL | Supplier will indicate YES / NO |
| Autosampler Temperature | At least in the range from 10°C – 40°C (at 20°C ambient temp. and rel. humidity < 25%) | Supplier will indicate YES / NO |
| Tubing material | PEEK, Radel | Supplier will indicate YES / NO |
| Measurement mode |  | Technical parameters offered by the supplier * |
| Detection principle | Time-correlated single photon counting of electrically actuated fluorescent DNA levers | Supplier will indicate YES / NO |
| Time-resolved switching at high frequencies | 1 kHz – 1 MHz, square wave | Supplier will indicate YES / NO |
| Low-frequency switching | 0.1 Hz – 4 Hz, square wave | Supplier will indicate YES / NO |
| Voltage sweep Staircase | ΔV > 1 mV | Supplier will indicate YES / NO |
| Constant voltage | ± 1 mV | Supplier will indicate YES / NO |
| Voltage range | ± 1 V | Supplier will indicate YES / NO |
| Sampling rate | 1 Hz – max. 5 Hz | Supplier will indicate YES / NO |
| Substrate material | Glass | Supplier will indicate YES / NO |
| Counter electrode | Indium tin oxide (ITO) | Supplier will indicate YES / NO |
| Regeneration |  | Technical parameters offered by the supplier * |
| Regeneration method | Oligonucleotide denaturation & hybridization | Supplier will indicate YES / NO |
| Standard conditions | Denaturation at pH 13-14, hybridization with 100 – 500 nM oligo | Supplier will indicate YES / NO |
| Regeneration cycle duration | Max. 8 min | Supplier will indicate YES / NO |
| Analyte molecules |  | Technical parameters offered by the supplier * |
| Proteins | Protein-protein Protein-DNA or Protein-RNA Protein-small molecule Protein-peptides | Supplier will indicate YES / NO |
| Antibodies | Antigen-Antibodies Bispecific analysis Biparatopic, triparatopic antibodies | Supplier will indicate YES / NO |
| Nucleic acids | DNA, RNA, LNA, PNA with single-nucleotide mismatch sensitivity DNA Origamis Structural RNA | Supplier will indicate YES / NO |
| Small molecules | DNA binding (intercalation) Protein-Small molecules | Supplier will indicate YES / NO |
| Chemical modifications and/or size | Proteins or nucleic acids which cause changes in conformation or charge | Supplier will indicate YES / NO |
| Cell-like structures | Virus-like particles, liposomes and lipid vesicles | Supplier will indicate YES / NO |

* Supplier will indicate YES / NO. If the supplier complements the minimum required technical parameters NE, this is the reason for excluding the tenderer from further participation 
in the procurement procedure.