## Technical Specification
Annex No. 1 - Technical Specifications
Type designation of the instrument
To be completed by contractor
Basic requirements of the contracting authority
It is required delivery of the stopped-flow spectrometer for the project CEITEC MU.
The required technical and functional characteristics | Requirement | Offer (applicants shall meet the required parameter verifiable manner, for example, specific reference to technical data sheets, drawings etc.) | Weight of the partial criterium in %
Dead-time of fast kinetic measurements lower than 2 ms in any configuration | yes |  | required
At constant temperature in range 5 – 60 °C with accuracy ± 0.1 °C (thermostat should be included) | yes |  | required
Measuring in extreme conditions (organic solvents, diluted acids) | yes |  | required
Measuring with low sample volumes (max. 60 µL total per experiment, 30 µL per reagent) | yes |  | required
Measuring in dual mixing mode (up to 3 samples) | yes |  | required
Measuring with variable volume mixing ratios (from 1:1 to 1:10) | yes |  | required
“Photomultiplier (PMT)” detection of fluorescence/absorbance | yes |  | required
Simultaneous detection of absorption and fluorescent analytical signal by PMT | yes |  | required
capability of measuring at anaerobic conditions | yes |  | required
“Quench-flow” module for monitoring fast reaction kinetics by other analytical methods | yes |  | required
Instrument controlled by PC (PC, monitor, printer, keyboard and mouse) with Microsoft Windows7 operating system | yes |  | required
Software for processing and analyzing 3D kinetic multidimensional experimental data | yes |  | required
Supplementary “stopped-flow” module for VARIAN Cary Bio 50 spectrometer | yes |  | required
Complementary instrumental set-up for measuring in lower (D2 from 200 nm) wavelength range | yes |  | required
Package of spare parts: xenon lamp, set of drive syringes: low-mid-high volume (2pcs each), pair of 1 ml stop syringes; set of O rings, tubing set,set of fuses | yes |  | required
Conductivity mode | advantage |  | 40 points
Supplementary “stopped-flow” module for any spectrofluorometer | advantage |  | 30 points
Simultaneous spectra acquisition in wavelength range min. 280 – 700 nm during the chemical reaction – “diode-array” detection | min. 280 - 700 nm |  | 240-700 nm….30 points, 250-700 nm….25 points, 260-700 nm….20 points, 270-700 nm….15 points, 280-700 nm....10 points