## Technicke_podminky
Annex No.1 - Technical specification Audio system for communication and stimulation in MRI Annex No.1 - Technical specification Audio system for communication and stimulation in MRI
Type designation of the device
To be completed by supplier
Basic requirements
Contracting authority require purchase 2 sets of the audio system for two-way communication with measured subjects and stimulation during acquisition in MRI machine.
Compulsory parameters
The system has to be fully compatible with SIEMENS Magnetom Prisma 3T and 64 channel head coil used in Multimodal and Functional Imaging Laboratory, CEITEC MU, placed in building A35 of University Campus Bohunice, Kamenice 5, 62500 Brno, Czech Republic. The system has to provide passive and active real-time noise cancelling. Two-way communication during scanning has to be possible in full duplex mode. Two-way communication between two scanners during scanning has to be possible. The system has to also allow recording the information about MR acquisition for temporal synchronization of recordings from the MRI and audio system. The headphones and microphone have to be constructed as in-bore devices compatible with MR acquisition sequences: T1 MPRAGE, T2, EPI BOLD, Multiecho and Multiband acquisition. Each set of the audio system has to consist of MR compatible headphones, MR compatible microphone and common control unit plus other necessary equipment. Detailed specification is placed below.
Technical and functional characteristics (The bids must meet all parameters set out below.) | Requested value | Supplier's offer (In case that the offered device fails to meet the technical requirements specified by the contracting authority, the tenderer state YES, otherwise NO.)
Headphones for MR scanner environment
Frequency Response | min. 60 -14,000 Hz | To be completed by supplier
Passive Noise Attenuation | min. 20 dB | To be completed by supplier
Total Harmonic Distortion | < 2% at 90 dB of maximal sound pressure level | To be completed by supplier
Active Noise Cancellation | min. 20 dB | To be completed by supplier
Total Noise Attenuation (combining passive + active) | min. 35 dB | To be completed by supplier
Maximum total thickness of earpads (with compressed foam) | <= 10 mm | To be completed by supplier
Connection between headphones and MRI control room via Optical Fiber Cable | Yes | To be completed by supplier
Length of optical cable | min. 12 m | To be completed by supplier
Compatible with Siemens Prisma 64-channel head-neck coil | Yes | To be completed by supplier
Microphone for MR scanner environment |  | To be completed by supplier
Frequency Response | min. 100-15,000 Hz | To be completed by supplier
Connection between microphone and MRI control room via Optical Fiber Cable | Yes | To be completed by supplier
Active noise reduction | Yes | To be completed by supplier
Accessories for mounting on head-coil (compatible with 64-channel head-neck coil and 20-channel head-neck coil for Siemens Prisma scanners) | Yes | To be completed by supplier
Total Harmonic Distortion | < 2% at 90 dB of maximal sound pressure level | To be completed by supplier
Control unit placed in MRI control room |  | To be completed by supplier
External audio input via SPDIF (optical) | Yes | To be completed by supplier
Necessary inputs and outputs for connection with headphones and microphone in MRI | Yes | To be completed by supplier
Synchronization with scanner via TTL BNC | Yes | To be completed by supplier
Built-in microphone | Yes | To be completed by supplier
Built-in loudspeaker | Yes | To be completed by supplier
Connection for communication between two units (two participants in different scanners) in full duplex mode including cables for physical connection (cable length 35 m) | Yes | To be completed by supplier
The seller (tenderer) declares that the delivery will comply with all the aforementioned purchaser requirements (contracting authority). If it becomes obvious in the course of preparating and implementing delivery that it is necessary to deliver additional devices, components or accessories or to perform any additional services or works in order to meet the purchaser's requirements hereunder, the seller hereby undertakes to deliver such devices and to perform such works or to provide such services as part of the original delivery under the agreement, with no increase of the purchase price (i.e. such deliveries, works or services shall not constitute extra deliveries or extra works).