## Technicke_podminky
Annex No.1 - Technical specification System for head movements and eye movements measurement
Type designation of the device
To be filled by supplier
Basic requirements
Contracting authority require purchase of the system for head movements and monocular eye tracking measurement during acquisition in MRI machine.
Compulsory parameters
The system has to be fully compatible with SIEMENS Magnetom Prisma 3T and 64 channel head coil used in Multimodal and Functional Imaging Laboratory, CEITEC MU, placed in building A35 of University Campus Bohunice, Kamenice 5, 62500 Brno, Czech Republic. Simultaneous recording of cameras for head movement evaluation and eye tracking complemented by visual stimulation via CRS BOLDscreen 24" LCD (not object of this tender) has to be possible. System has to also allow recording the information about MR acquisition for temporal synchronization of recordings from MRI and tracking system. Cameras have to be constructed as a in-bore devices compatible with MR acquisition sequences: T1 MPRAGE, T2, EPI BOLD, Multiecho and Multiband acquisition with gradient-echo and spin-echo techniques. The camera for head movement measurement has to have HD resolution. Detailed specification is placed below.
Technical and functional characteristics (The bids must meet all parameters set out below.) | Requested value | Supplier's offer (In case that the offered device fails to meet the technical requirements specified by the contracting authority, the tenderer state YES, otherwise NO.)
Certificate of MR compatibility for cameras and accessories placed in MR shielded room with 3T magnetom Siemens Prisma | Yes | To be filled by supplier
Compatibility of cameras and accessories with 64 channel head coil for 3T Siemens Magnetom Prisma | Yes | To be filled by supplier
Possibility of visual stimulation of the participant via LCD monitor and mirror has to remain | Yes | To be filled by supplier
Possibility of simultaneous usage of eye tracking, head movement system and visual stimulation through the mirror placed on the head coil. | Yes | To be filled by supplier
Possibility of separate usage of head movement tracking system | Yes | To be filled by supplier
Possibility of separate usage of eye tracking system | Yes | To be filled by supplier
Wiring required for using of all devices | Yes | To be filled by supplier
Ability to receive and place sychronization pulses from MR scanner into the measured data | Yes | To be filled by supplier
Warranty | min. 2 years | To be filled by supplier
Eye tracking - monocular | Yes | To be filled by supplier
Eye tracking - sampling rate | min. 50 Hz | To be filled by supplier
Eye tracking - software for recording and data analysis | Yes | To be filled by supplier
Eye tracking - in-bore applicable camera | Yes | To be filled by supplier
Eye tracking - possibility of hardware upgrade to binocular systém in the future | Yes | To be filled by supplier
Head movement - resolution of the camera | min. 1280 x 720 | To be filled by supplier
Head movement - sampling rate of the camera | min. 50 Hz | To be filled by supplier
Head movement - software for data recording | Yes | To be filled by supplier
Head movement - in-bore applicable camera | Yes | To be filled by supplier
Device reference from other research facility | Yes, min. 1 | To be filled by supplier
The seller (tenderer) declares that the delivery will comply with all the aforementioned requirements of the purchaser (contracting authority). If it becomes obvious in the course of the delivery preparation and implementation that it is necessary to deliver additional devices, components or accessories or to perform any additional services or works in order to meet the purchaser's requirements hereunder, the seller hereby undertakes to deliver such devices and to perform such works or to provide such services as a part of the original delivery under the agreement, without any increase of the purchase price (i.e. such deliveries, works or services shall not constitute extra deliveries or extra works).