## Technicke podminky
Annex No. 1 - Technical Specifications Side-entry cryo-electron microscopy specimen holder for Talos F200C transmission electron microscope
Typové označení přístroje / Type of device
Vyplní dodavatel/Fill in by a tenderer
Základní požadavky zadavatele / Basic requirements
The tender comprises purchase of the side-entry sample holder for the ThermoScientific Talos F200C transmission electron microscope for cryo-electron microscopy applications, and the training in optimal utilization of the holder.
Required technical and functional qualities | Minimal technical requirements | Nabídka uchazeče Pokud je zadavatelem u daného parametru požadován číselný údaj, je ho uchazeč povinen uvést, v opačném případě bude zadavatel vycházet z jím minimální stanovené hodnoty - viz. sloupec Požadovaná hodnota, uchazeči uvedou splnění požadovaného parametru ověřitelným způsobem, např. konkrétním odkazem na technické listy, výkresy apod. / Tenderer indicates meeting all the technical and functional requirements; shall be evidenced unambiguously and plausibly, e.g. by providing technical sheets and drawings of the manufacturer.
The holder accepts at least 2 TEM grids in the ThermoScientific Autogrid. | YES | Vyplní dodavatel/Fill in by a tenderer
The tilt range inside the microscope is at least -64° - +64°. | YES | Vyplní dodavatel/Fill in by a tenderer
The tip temperature inside the microscope is < -170°C with the temperature stability better than +/- 0.1°C per hour. | YES | Vyplní dodavatel/Fill in by a tenderer
The holder operation time at the temperatures below -165°C without refilling LN2 is at least 8 hours. | YES | Vyplní dodavatel/Fill in by a tenderer
The holder dewar has centrosymmetric design. | YES | Vyplní dodavatel/Fill in by a tenderer
Drift rate at 0° stage tilt 30min after insertion into the TEM < 2nm/min. | YES | Vyplní dodavatel/Fill in by a tenderer
Isotropic resolution at 0° stage tilt <2.5Å. | YES | Vyplní dodavatel/Fill in by a tenderer
Workstation for sample loading with LN2 hold time of >30min. | YES | Vyplní dodavatel/Fill in by a tenderer
Temperature controller and temperature monitor. | YES | Vyplní dodavatel/Fill in by a tenderer
Pumping station or adaptors for the attachment of the holder to Gatan Solarus II or Gatan Model 655 Turbo pumping station. | YES | Vyplní dodavatel/Fill in by a tenderer
The seller (tenderer) declares that the delivery will comply with all the aforementioned requirements of the purchaser (contracting authority). If it becomes obvious in the course of the delivery preparation and implementation that it is necessary to deliver additional devices, components or accessories or to perform any additional services or works in order to meet the purchaser's requirements hereunder, the seller hereby undertakes to deliver such devices and to perform such works or to provide such services as a part of the original delivery under the agreement, without any increase of the purchase price (i.e. such deliveries, works or services shall not constitute extra deliveries or extra works).