Technical conditions and technical specification 		Appendix No. 1

System for synchronous recording of Ca2+ transients and contraction in cardiomyocytes

Number of the system:	1

Annotation - General requirements of the submitter on the equipment:

The system has to enable synchronous recording of Ca2+ transients and contraction in cardiomyocytes, both freshly isolated and hiPSC-derived ones. The system should be composed of the following key components:

- inverted fluorescent microscope with 4UV-excitation

- field stimulator

- stimulation chamber with exchange of the perfusion solution and with the system for monitoring and control of the temperature of the perfusion solution at least in the range between 30 and 45°C including stimulation electrodes

- perfusion cannula for exchange of the solution in the near surroundings of the measured cell (flow directed to the measured cell with a cannula), with control of the temperature of the perfusion solution, at least in the range between 30 and 45°C (temperature controller included)

- micromanipulator for the perfusion cannula including magnetic stand and steel base plate

- fluorescence-excitation light source with fast switching of the wave lengths for dyes with double excitation

- fluorescence recording system

- camera

- A/D converter

- PC for recording and analysis of the data, including compatible LCD monitor (at least 24", resolution at least 1920 x 1080 pixels), keyboard, and mouse

- software for recording and analysis of the data (Ca2+ transients and contraction in cardiomyocytes, both using the sarcomere length in adult cardiomyocytes and another system applicable in ‘amorphous’ hiPSC-derived cardiomyocytes) - sampling rate at least up to 1 kHz

- rotatable emission shutter for limiting/controlling background fluorescence

All components have to be fully compatible allowing smooth running of the above-mentioned recordings and analysis of the acquired data.

230 V, 50-60 Hz

Key technical requirements of some of the components:

- Inverted fluorescent microscope

Number of items: 1 pc

| Model – type/ product designation | Manufacturer |
| --- | --- |

| Parameter | TECHNICAL CONDITIONS required | TECHNICAL SPECIFICATION offered |
| --- | --- | --- |
| fluorescence-excitation light source | LED, at least 3 filters for different wave lengths, 340 nm/385 nm, and 485 nm |  |
| objectives | 40x |  |
| colour corrected infinity optical system | YES |  |
| stage | mechanical, long wand |  |
| head | trinocular |  |
| stand | wide, stable |  |

- Field stimulator

Number of items: 1 pc

| Model – type/ product designation | Manufacturer |
| --- | --- |

| Parameter | TECHNICAL CONDITIONS required | TECHNICAL SPECIFICATION offered |
| --- | --- | --- |
| voltage stimulation | YES, at least in the range +/- 40 V |  |
| impulses - character | bipolar, unipolar (positive, negative) |  |
| impulses – duration | at least in the range between 0.5 – 80 ms |  |
| impulses – frequency | at least in the range between 0.04 – 90 Hz |  |

- Camera

Number of items: 1 pc

| Model – type/ product designation | Manufacturer |
| --- | --- |

| Parameter | TECHNICAL CONDITIONS required | TECHNICAL SPECIFICATION offered |
| --- | --- | --- |
| compatibility with the microscope and PC+software | YES |  |
| sensor resolution | at least 782 x 778 pixels |  |
| adjustable frame rate value | YES, the maximal rate at least 250 Hz |  |

- PC with Software for recording and analysis of the data

Number of items: 1 pc

| Model – type/ product designation |  | Manufacturer |  |
| --- | --- | --- | --- |
| Parameter | TECHNICAL CONDITIONS required |  | TECHNICAL SPECIFICATION offered |
| RAM memory | at least 16 GB |  |  |
| SSD disc | at least 512 GB |  |  |
| USB port | at least 5x in total, at least two of them USB 3.0 |  |  |
| Windows | YES, at least Windows 10 |  |  |
| Software for recording and analysis of the data |  |  |  |
| Software name and manufacturer |  | SW function and feature description |  |

- Monitor

Number of items: 1 pc

| Model – type/ product designation |  | Manufacturer |  |
| --- | --- | --- | --- |
| parameter | TECHNICAL CONDITIONS required |  | TECHNICAL SPECIFICATION offered |
| diagonal | at least 24" |  |  |
| resolution | at least 1920 x 1080 pixels |  |  |

|  | these fields marked in yellow are filled in by the supplier |
| --- | --- |