Technické podmínky

Technické podmínky jsou vyhotoveny jak v českém, tak anglickém jazyce. V případě rozporů mezi českou 
a anglickou jazykovou verzí, nebo v případě sporu o výklad použitých pojmů, má přednost česká jazyková verze.

Peptidový syntetizátor

Popis zařízení a jeho využití

Peptidový syntetizátor pro plně automatizovanou syntézu peptidů na pevné fázi (SPPS). Přístroj umožní paralelní syntézu několika peptidů od prvního couplingu až po odštěpení peptidu z pevné fáze. Zahřívání reakční cely umožní zrychlenou syntézu peptidů a vyšší výtěžky surového produktu u náročnějších sekvencí.

Předmětem dodávky je kompletní, nové, nepoužité a zcela funkční zařízení.

| Parametr | Parametr nabízený dodavatelem |
| --- | --- |
| Výrobce |  |
| Typ/Model |  |

| Minimální požadované technické parametry | Technické parametry nabízené dodavatelem* |
| --- | --- |
| Možnost syntézy alespoň šesti peptidů zároveň (paralelně, nikoliv sekvenčně) v uzavřených reakčních celách | (Dodavatel uvede ANO/ NE a skutečnou hodnotu) |
| Zahřívání reakčních cel (pomocí mikrovln, indukce nebo infračerveného záření) | (Dodavatel uvede ANO/ NE) |
| Možnost používat opakovaně použitelné cely ze skla nebo odolného plastu o různých objemech, které jsou vhodné na různá množství syntetizovaného peptidu | (Dodavatel uvede ANO/ NE) |
| Minimální rozsah množství syntetizovaného peptidu mezi 5 µmol–1 mmol v každé reakční cele | (Dodavatel uvede ANO/ NE a skutečnou hodnotu) |
| Fluidní systém zabraňující křížové kontaminaci mezi jednotlivými potrubími a reakčními celami | (Dodavatel uvede ANO/ NE) |
| Možnost připojit alespoň 8 nádob na rozpouštědla a alespoň 27 nádob na aminokyseliny | (Dodavatel uvede ANO/ NE a skutečnou hodnotu) |
| Možnost pre-aktivace reagentů | (Dodavatel uvede ANO/ NE) |
| Možnost přidání celého obsahu nádoby na aminokyseliny do vybraných reakčních cel bez nutnosti primingu, aby se minimalizovalo plýtvání drahými neproteinogenními aminokyselinami nebo jinými stavebními bloky | (Dodavatel uvede ANO/ NE) |
| Časované přídavky nebo měření pevnou smyčkou („fixed-loop“) pro přídavek přesných objemů kapalin | (Dodavatel uvede ANO/ NE) |
| Možnost volby způsobu míchání obsahu reakční cely: probublávání dusíkem nebo třepání | (Dodavatel uvede ANO/ NE) |
| Plně automatizované a programovatelné štěpení peptidů z pevné fáze přímo v přístroji bez manuálního zásahu | (Dodavatel uvede ANO/ NE) |
| Robustní potrubí odolávající kyselině trifluoroctové | (Dodavatel uvede ANO/ NE) |
| Řídící software na specializovaném počítači | (Dodavatel uvede ANO/ NE) |
| Bezpečnostní prvky včetně automatického ukončení sekvence a promytí pevné fáze v případě chyby, e-mailové nebo SMS upozornění v případě chyby | (Dodavatel uvede ANO/ NE) |
| Startovací sada obsahující alespoň všech 20 proteinogenních aminokyselin (Fmoc-AA-OH s ochráněnými postranními řetězci; min. 25 mmol), pevnou fázi s Rink amid linkerem (0,1 mmol), pevnou fázi s Gly-modifikovaným Wangovým linkerem (0,1 mmol), aktivátor HCTU (100 g), všechny potřebné cely a nádoby, vzorky pro kontrolu kvality | (Dodavatel uvede ANO/ NE) |
| Minimální požadavky na možnost budoucího rozšíření přístroje o další moduly a funkce: |  |
| UV monitorování reakčních roztoků v každé cele v reálném čase pro optimalizaci kroků odchraňování | (Dodavatel uvede ANO/ NE) |

*Dodavatel uvede ANO/NE a doplní požadované informace. Pokud dodavatel doplní do Minimálních požadovaných technických parametrů NE, je to důvod pro vyloučení uchazeče z další účasti v zadávacím řízení. Dodavatel je povinen přiložit k této technické specifikaci i svou vlastní technickou specifikaci či svůj vlastní popis zařízení.

The technical conditions

The technical conditions are in both Czech and English. In the event of a discrepancy between the Czech and English language versions, or in the event of a dispute as to the interpretation of the terms used, the Czech language version shall prevail.

Peptide synthesizer

Equipment description and its use

A peptide synthesizer for fully automated solid-phase peptide synthesis (SPPS). The instrument will enable the parallel synthesis of multiple peptides from the first coupling to peptide cleavage off the resin. Heating the reaction vessel will allow for expedited peptide synthesis and high crude yields of challenging sequences.

The delivery includes complete, new, unused and fully functional equipment.

| Parameter | Parameter offered by supplier |
| --- | --- |
| Producer |  |
| Typ/Model |  |

| Minimal required technical parameters | Technical parameters offered by the supplier* |
| --- | --- |
| Ability to synthesize at least six peptides at the same time (parallel, not consecutive) in closed reaction vessels | (Supplier shall indicate YES/ NO and actual value) |
| Reaction vessel heating (via microwave, induction, or infra-red light) | (Supplier shall indicate YES/ NO) |
| The possibility to use reusable glass or durable plastic vessels of different volumes appropriate for the synthesis scale | (Supplier shall indicate YES/ NO) |
| Minimum scale range of 5 µmol–1 mmol in each reaction vessel | (Supplier shall indicate YES/ NO and actual value) |
| A fluidic system preventing cross-contamination between individual lines and vessels | (Supplier shall indicate YES/ NO) |
| The possibility to connect at least 8 solvent containers and at least 27 amino acid containers | (Supplier shall indicate YES/ NO and actual value) |
| The possibility of reagent pre-activation | (Supplier shall indicate YES/ NO) |
| The possibility of adding the entire contents of amino acid containers to selected reaction vessels without priming to minimize the waste of expensive non-native amino acids or other building blocks | (Supplier shall indicate YES/ NO) |
| Timed deliveries or fixed loop measurements to deliver precise liquid volumes | (Supplier shall indicate YES/ NO) |
| The possibility to select the agitation method of reaction vessel contents: nitrogen bubbling or shaking | (Supplier shall indicate YES/ NO) |
| Fully automated and programmable on-board cleavage of peptides off the resin without manual intervention | (Supplier shall indicate YES/ NO) |
| Robust tubing able to withstand trifluoroacetic acid | (Supplier shall indicate YES/ NO) |
| Control software on a dedicated computer | (Supplier shall indicate YES/ NO) |
| Safety features including automated sequence termination and resin rinsing in case of an error, email or SMS notifications in case of an error | (Supplier shall indicate YES/ NO) |
| Starter kit containing at least all 20 canonical amino acids (Fmoc-AA-OH with side-chain protection; min. 25 mmol), Rink amide resin (0.1 mmol), Gly-loaded Wang resin (0.1 mmol), HCTU activator (100 g), all required vessels and vials, and quality control samples | (Supplier shall indicate YES/ NO) |
| Minimum requirements for future expansion of the device with additional modules and functions: |  |
| Real-time UV monitoring of reaction solutions in each vessel for optimization of deprotection steps | (Supplier shall indicate YES/ NO) |

*The supplier shall indicate YES/NO and complete the required information. If the Supplier adds NO to the Minimum Required Technical Parameters, this shall be grounds for disqualification of the Bidder from further participation 
in the procurement process. The Supplier shall be required to attach its own technical specification or its own description of the equipment to this Technical Specification.